Crystal Structure Report for Alex Jen

by Dylan Rogers & Werner Kaminsky

Sample Submitted by Francis Lin 

ID – BoSeRB_4F

X-ray Crystallography Laboratory

University of Washington

08. August 2019
Data Collection  started: 6/24/2019
Documentation finished: 8/8/2019
Data backup accessibly through http:

URL: http://cad4.cpac.washington.edu/structures
W.Kaminsky: tel. (206) 543 7585

(x-ray lab):     tel. (206) 543 0210 

Email: kaminsky@chem.washington.edu
Understanding crystallographic information is not trivial. If you discuss a structure in a publication, please let the crystallographers proof-read the final draft of that manuscript before you submit it.

For many structures it is true that crystallography is an intellectual performance (like synthesis or elucidation of mechanisms) and not a routine job. In these cases, particularly when the structures contribute significantly to the content of the publication, the crystallographers should be co-authors. Do not think this is not necessary because you pay for structures. 

A black piece, measuring 0.11 x 0.15 x 0.32 mm3 was mounted on a loop with oil.  Data was collected at -173oC on a Bruker APEX II single crystal X-ray diffractometer, Mo-radiation.
Crystal-to-detector distance was 40 mm and exposure time was 60 seconds per frame for all sets.  The scan width was 1o.   Data collection was 94.3% complete to 25o  in (.  A total of 31724 unique reflections were collected covering the indices, -13<=h<=13, -22<=k<=22, -30<=l<=30.  with Rint = 0.2200.  Indexing and unit cell refinement indicated a triclinic lattice.  The space group was found to be P(1 (No. 2). 

The data was integrated and scaled using SAINT, SADABS within the APEX2 software package by Bruker.
 

The contribution of disordered chloroform to the diffraction pattern was removed with SQUEEZE

Solution by direct methods (SHELXT
 or SIR97
) produced a complete heavy atom phasing model consistent with the proposed structure. The structure was completed by difference Fourier synthesis with SHELXL.
,
 Scattering factors are from Waasmair and Kirfel
. Hydrogen atoms were placed in geometrically idealised positions and constrained to ride on their parent atoms with C---H distances in the range 0.95-1.00 Angstrom. Isotropic thermal parameters Ueq were fixed such that they were 1.2Ueq of their parent atom Ueq for CH's and 1.5Ueq of their parent atom Ueq in case of methyl groups. All non-hydrogen atoms were refined anisotropically by full-matrix least-squares.
The very long side chains become increasingly delocalized with distance to the backbone which caused the refinement to not completely converge.
The structure is of marginal quality and may be added for completion of data only to a publication. Table 1 summarizes the data collection details. Figure 1 shows an ORTEP
 of the asymmetric unit.
[image: image1.jpg]



Figure 1. ORTEP of the structure with thermal ellipsoids at the five % probability level. 
Table 1: Crystallographic data for the structures provided.

Empirical formula 
C90 H102 F4 N8 O2 S3 Se2

Formula weight 
1657.89

Temperature 
100(2) K

Wavelength 
0.71073 Å

Crystal system 
Triclinic

Space group 
P -1

Unit cell dimensions
a = 13.6204(13) Å
= 70.177(7)°.


b = 23.304(3) Å
= 83.982(6)°.


c = 31.303(3) Å
 = 80.737(6)°.

Volume
9211.8(18) Å3
Z
4

Density (calculated)
1.195 Mg/m3
Absorption coefficient
0.927 mm-1
F(000)
3464

Crystal size
0.320 x 0.150 x 0.110 mm3
Theta range for data collection
1.385 to 25.821°.

Index ranges
-13<=h<=13, -22<=k<=22, -30<=l<=30

Reflections collected
31724

Independent reflections
31724 [R(int) = 0.2200]

Completeness to theta = 25.000°
94.3 % 

Refinement method
Full-matrix least-squares on F2
Data / restraints / parameters
31724 / 2768 / 1951

Goodness-of-fit on F2
1.300

Final R indices [I>2sigma(I)]
R1 = 0.2238, wR2 = 0.4797

R indices (all data)
R1 = 0.4724, wR2 = 0.5791

Largest diff. peak and hole
1.526 and -0.808 e.Å-3
� Bruker (2007) APEX2 (Version 2.1-4), SAINT (version 7.34A), SADABS (version 2007/4), BrukerAXS Inc, Madison, Wisconsin, USA.


� (a) A. Spek, J. Appl. Cryst. (2003), 36, 7-13; 


   (b) P. van der Sluis & A. L. Spek, Acta Cryst. (1990). A46, 194-201.


   (c) Spek, A. L. (2009). Acta. Cryst. D65, 148-155.


� Sheldrick GM. (2015) SHELXT - Integrated space-group and crystal-structure determination. Acta Cryst. A71, 3-8.


� (a) Altomare A, Burla C, Camalli M, Cascarano G L, Giacovazzo C, Guagliardi A,   Moliterni AGG, Polidori G, Spagna R. (1999) SIR97: a new tool for crystal structure determination and refinement Journal of Applied Crystallography, 32, 115-119. 


   (b) Altomare A, Cascarano G L, Giacovazzo C, Guagliardi A. (1993) Completion and refinement of crystal structures with SIR 92. Journal of Applied Crystallography, 26, 343-350.


� (a) Sheldrick GM. (1997) SHELXL-97, Program for the Refinement of Crystal Structures. University of Göttingen, Germany.


   (b) Sheldrick GM. (2015) Crystal structure refinement with SHELXL. Acta Cryst. (2015). C71, 3–8


� Mackay, S.; Edwards, C.; Henderson, A.; Gilmore, C.; Stewart, N.; Shankland, K.; Donald, A. (1997) MaXus: a computer program for the solution and refinement of crystal structures from diffraction data. University of Glasgow, Scotland,.


� Waasmaier, D.; Kirfel, A. (1995) New Analytical Scattering Factor Functions for Free Atoms and Ions. Acta Crystallographica A., 51, 416-430.


� Farrugia LJ. (1997) Ortep-3 for Windows. Journal of Applied Crystallography, 30, 565





