Crystal Structure Report for J Kovacs

Submitted by Erica Shaffer 

Sample ID – shaffer003

by R Swartz  &W Kaminsky

X-ray Crystallography Laboratory

University of Washington

07. July 2008
Data backup accessibly through http:

URL: 
http://cad4.cpac.washington.edu/structures
Username: 
xray-user

Password: 
hkl-data

W.Kaminsky: tel. (206) 543 7585

(x-ray lab):     tel. (206) 543 0210 

Email: kaminsky@wintensor.com
Understanding crystallographic information is not trivial. If you discuss a structure in a publication, please let the crystallographers proof-read the final draft of that manuscript before you submit it.

For many structures it is true that crystallography is an intellectual performance (like synthesis or elucidation of mechanisms) and not a routine job. In these cases, particularly when the structures contribute significantly to the content of the publication, the crystallographers should be co-authors. Do not think this is not necessary because you pay for structures. Even I  have to pay for my own structures for my research.

A yellow prism, 0.59 x 0.53 x 0.29 mm was mounted on a glass capillary with oil.  Data was collected at -143oC.

Crystal-to-detector distance was 30 mm and exposure time was 30 seconds per degree for all sets.  The scan width was 1.1o.   Data collection was 98.6% complete to 25o  in (.  A total of  146639 partial and complete reflections were collected covering the indices, h = -13 to 13, k = -19 to 19,  l = -24 to 25.  13618 reflections were symmetry independent and the Rint = 0.0932 indicated that the data was of slightly less than average quality (0.07). 

Indexing and unit cell refinement indicated a triclinic P lattice.  The space group was found to be P(1 (No.2). 

The data was integrated and scaled using hkl-SCALEPACK. This program applies a multiplicative correction factor (S) to the observed intensities (I) and has the following form:


[image: image1.wmf]22

2(sin)/

S = ()/scale

B

e

ql

-


S is calculated from the scale and the B factor determined for each frame and is then applied to I to give the corrected intensity (Icorr).

Solution by direct methods (SIR97) produced a complete heavy atom phasing model consistent with the proposed structure. 

All hydrogen atoms were located using a riding model. All non-hydrogen atoms were refined anisotropically by full-matrix least-squares. 

A solvent molecule, CH2OH  was found,  that occupies only 42% of the unit cell. The presence of the solvent causes one of the tetraphenylboranes inside the asymmetric unit to be disordered as well. This disordered tetraphenylborane is filling available space as good as possible but with the problem of keeping the phenyls apart, putting considerable strain on the phenyl containing H77A and causing almost a situation of collision with H18 of the main complex. The Mn-dimer is bridged by a sulphur and by the oxygen atoms of a tetrafluoroacetate that shows signs of strong librative motion of the fluorines.

This challenging structure was edited for publication and is of high quality.




_1097643391.unknown

