Structure Report for Sasaki

Crystal submitted by Byung Jn Kim 

Sample ID – uobj5
Structure by Werner Kaminsky

X-ray Crystallography Laboratory

University of Washington

26. September 2011
Data backup accessibly through http:

URL: 
http://cad4.cpac.washington.edu/structures
Username: 
xray-user

Password: 
hkl-data

W.Kaminsky: tel. (206) 543 7585

(x-ray lab):     tel. (206) 543 0210 

Email: kaminsky@wintensor.com
Understanding crystallographic information is not trivial. If you discuss a structure in a publication, please let the crystallographers proof-read the final draft of that manuscript before you submit it.

For many structures it is true that crystallography is an intellectual performance (like synthesis or elucidation of mechanisms) and not a routine job. In these cases, particularly when the structures contribute significantly to the content of the publication, the crystallographers should be co-authors. Do not think this is not necessary because you pay for structures. Even I  have to pay for my own structures for my research.

A colorless crystal plate, 0.20 x 0.18 x 0.07 mm3 of size, was mounted on a glass capillary with oil.  Data was collected at -143oC.
Crystal-to-detector distance was 40 mm and exposure time was 70 seconds per degree for all sets.  The scan width was 2o.   Data collection was  96.8% complete to 25o  in (.  A total of  32467 partial and complete reflections were collected covering the indices, h = -15 to 15, k = -11 to 11,  l = -36 to 36.  7280 reflections were symmetry independent and the Rint = 0.1995 indicated that the data was of less than average quality (0.07). 

Indexing and unit cell refinement indicated a monoclinic P lattice.  The space group was found to be P 21 (No.4). 

The data was integrated and scaled using hkl-SCALEPACK. This program applies a multiplicative correction factor (S) to the observed intensities (I) and has the following form:


[image: image1.wmf]22

2(sin)/

S = ()/scale

B

e

ql

-


S is calculated from the scale and the B factor determined for each frame and is then applied to I to give the corrected intensity (Icorr).

There was no direct solution to this structure (SIR97, SIR92, DIRDIFF, SHELXS97 all failed). However, because of a strong similarity to uobj4, the previously solved structure was successfully applied as a starting point for solving the phase problem. The structure was then corrected and refined with shelxl97.
All hydrogen atoms were located using a riding model. All non-hydrogen atoms were refined anisotropically by full-matrix least-squares.



_1097643391.unknown

